Key things:  We should focus on functional decomposition and modularity of our system.

Dr Rob would prefer emails for definite and burning questions.

Dr. Rob would prefer GitHub

Introductory Background: Dr. Read has previously completed seven similar projects for Rice University. The project’s purposes:

- Help Kids OR create safer pediatric operating rooms (ORs).

- Support the students’ (Design Team) educational and professional goals.

Public Invention can help provide continuity after the current design team completes work, this is why the project was chosen to be open source.

- An anesthesia machine has an open circuit system, where excess and waste gases can leave. This is collected by the scavenging system and is carried out of the OR.

Dr. Read was thinking of connecting the machine to tubing that conducts gas outside of the container? Perhaps with a small fan. Consider using a sensor, connected to a microcontroller, should be automatic and require little attention from medical staff.

Kids OR Environment

- Kids OR primarily operates in Africa, and low-resource regions all around the world.

- An OR is made out of a shipping container, with art and other materials to make the environment more cozy for children

- A facility may perform approximately 10-20 pediatric surgeries daily, both routine and more complex procedures.

- The anesthesia machines used are not different to those in the US, but the issue comes from a lack of ventilation and built-in anesthetic gas scavenging systems.

- Many Kids ORs have independent solar-panels. Electricity is available but is inconsistent and unreliable, as they are often not connected to the local electric grid (which may also be unstable).

Current Problem

- Only 1 in 10 Kids OR rooms have a gas scavenging system. Throughout a day of surgery, medical personnel report drowsiness and wooziness, suggesting exposure to waste anesthetic gas

- Typically, the child inhales the anesthesia and exhales gas that still contains some anesthetic. The problem? Without a gas scavenging system, these small amounts of WAGs are being released into the OR directly.

- We are limited by space, electrical power, ventilation inside container, and tripping hazards

Constraints/Challenges

- In anesthesia machines, there are only partially sealed suction fans. This partial sealing is to prevent excess pressure on the breathing circuit and lungs. Our solution cannot use excess or too powerful suction, or we risk patient endangerment.

- Chemically scrubbing or destroying the anesthetic gas may not be the main priority or the simplest way to combat the problem.

- According to Dr. Robert, the environmental effects of gas release outdoors is minimal compared to the exposure risk in the OR. This is something the team will further verify and explore!

- The tubing to carry the waste gas out of OR can be a complicated part of the design problem! We must avoid causing any serious tripping hazard in the OR, and given that the Kids OR is made out of a shipping container, we may not be able to build the tubing into the floor, wall, or ceiling (as is typically done). A solution cannot cause a major change to the OR.

- Ensure modularity: This is so if the system breaks, repair is simple and easy, as well as design improvements!

- Low cost, accessibility, reliability, and compatibility

Potential Questions for Kids OR Team

What parameters control the required gas flow rate through the scavenging system?

What is the typical flow rate of waste gas leaving the anesthesia machine?

How should the system respond if adequate airflow is not present or waste gas is detected in the room (i.e. what is the ideal response)?

Are there anesthesiologists, surgeons, or other medical personnel who have worked inside KidsOR facilities and could be interviewed about the workflow and exposure problem?

What do the rooms for Kids OR look like?

Is there a preventative maintenance procedure already being followed for this?

What type of anesthetic gas is being used?